## Preliminary Test Guideline
HbA1c (NGSP) System Calibration and Preliminary Test Guideline
ASRL#1
[Reference Material Institute for Clinical Chemistry Standards (ReCCS), Japan]
Purpose:
System Calibration and Preliminary Test of HbA1c (NGSP) Measurement System
Procedure:
Please, measure the HbA1c % in the calibrator JCCAL-SC specimens using your measurement method and instrument.
(Type A) Hemolysed erythrocytes in a physiological solution (not including plasma). Total Hb concentration is approximately 12~15g/dL.
(Type B) Separated and concentrated red cells (not including plasma). Total Hb concentration is approximately 30g/dL.
(Type C) Hemolysed whole blood. Total Hb concentration is approximately 12~15g/dL.
1. The specimens will be shipped to you by ASRL#1 (Laboratory of ReCCS, Japan) frozen and on a dry ice.
2. Upon receipt, please check if some dry ice still remains in the shipping box, and immediately store the specimens
in a deep freezer (-70℃ or below) until measurement.
3. Each specimen should be analyzed two times for preliminary test in one assay date (Result 1, Result 2).
4. After thawed at room temperature, please mix several times the content of each vial,
and measure within 2 hours HbA1c values by appropriate diluting of the samples
manually or automatically as explained in your instrument procedure manual.
Mesurement Results Report:
Please send the Preliminary Test Measurement Results using the Report Form to ASRL#1 (ReCCS)
at the following e-mail address:
E-mail: ngspdata@reccs.net
Deadline： When you finish the preliminary test, please report your results as soon as you can.
Preliminary test is should be done before the certification test upon sample arrival.
Refer to "Summary of NGSP Certification_ASRL#1".
Tel (general inquiry and orders): 81-45-507-6145 Fax: 81-45-530-9036
Tel (technical inquiry): +81-45-507-6326 Fax: +81-45-507-7197

## Preliminary Test Report Form
Preliminary Test Measurement Results Report
HbA1c(NGSP) Preliminary Test
HbA1c(NGSP) Measurement Results （NGSP %）
JCCAL-SC4 Specimen ID | Day 1
| Result 1 | Result 2
SC4-1
SC4-2
SC4-3
SC4-4
SC4-5
SC4-6
SC4-7
SC4-8
SC4-9
SC4-10
Assay Date
Instrument
Reagent ID & Lot # (e.g. buffers, hemolyzing reagents, reagent cartridges, etc)
Reference Material
Manufacturer's calibrator
Sample Dilution Buffer
Dilution Ratio
Dilution Method
Injected Sample Volume
Laboratory Name
Department
Address
Operating Person Adviser
Operating Person
Telephone Number
FAX Number
E-mail Address
Preliminary test is required before the certification test! It should be done upon sample arrival.
Refer to "Summary of NGSP Certification_ASRL#1".
Please send the Preliminary Test Measurement Results using the Report Form to ASRL#1 (ReCCS)
at the following e-mail address:
E-mail: ngspdata@reccs.net
Deadline：When you finish the preliminary test,
please report your results as soon as you can.